EQUIPMENT MANAGEMENT PROCEDURE

1.0 	 Purpose

The purpose of this procedure is to ensure that testing and calibration equipment used in the laboratories are documented and information on these pieces of equipment is readily available to staff who might need to use the equipment; that such equipment are regularly serviced, calibrated; and that the software of computer-based equipment is validated and subject to the defined change/version control system.

- Scope

  - This procedure applies to all testing equipment used in the laboratories including physical standards.

  - This procedure describes how details relating to specific items of test are documented and sets out the system for documenting all calibrations and maintenance performed.

  - It also defines the system of controls applied to software of computer-based equipment or configuration management.

- References:

  - AQC.QM-001

  - ISO/IEC 17025 Clause 5.5

  - EA-4/07, Traceability of measuring and test equipment to National Standards

  - NIS 37, A guide to managing the configuration of computer systems (Hardware, software and firmware) Used in NAMAS accredited laboratories

  - ENGA (98) 95, Guidelines for the use of computers and computer systems in accredited laboratories

4.0 ABBREVIATIONS/TERMS USED:

- The terms used in AQC.QM-001and AQC.QMS.SOP-02shall apply.

- ERF: Equipment Record File

- CMR: Configuration Management Record

5.0 RECORDS OF CHANGE

| NO. | Date | Details of changes |  | Title/ Name | Signature |
| --- | --- | --- | --- | --- | --- |
|  | (dd-mm-yyyy) | Page | Clause/Sub-clause and comment | By: |  |
| 01 | 30/06/2021 | All | • Included Tables and Appendices on the Header note. • Changed Version No. from 00 to 01 | QM |  |

6.0 DISTRIBUTION TABLE

The controlled copies of this document are issued in accordance with appendix D.

| DOC. COPY NO. | DOC. VERSION |  |
| --- | --- | --- |
| 01 | 01 | QUALITY MANAGER |
| 02 | 01 | TECHNICAL MANAGER |
| 03 | 01 | ACCREDITATION BODY (KENAS) |

7.0	Principal Responsibilities

  - The Technical Manager and Quality Manager are responsible for ensuring that the requirements of this procedure are implemented for all relevant equipment and reference materials in their area.

  - All staff conducting tests and measurements is responsible for ensuring that equipment and reference materials are used within the constraints described in each equipment and reference materials records.

- Procedures

    - An inventory of all items of equipment significant to laboratory for test shall be given. This inventory shall consist of the equipment description, the unique equipment inventory number and equipment serial numbers or model or type where available and also calibration status (see Appendix C).

    - The laboratory Equipment referred to in this case includes: general service equipment; volumetric equipment; measuring equipment or instruments.

    - Details of each item of equipment (given in the inventory mentioned in 8.1.1) shall be entered onto an Equipment Record File (ERF) and indexed by' unique inventory number of the equipment. Details concerning the generation of equipment inventory numbers are given in Appendix A.

  - New Equipment Commissioning

    - All new equipment will be inspected as per procedure for acceptance against laboratory requirements by the analyst.

    - Once accepted, the analyst or deputy will commission the equipment according to the manufacturer's instructions or will organize a manufacturer's service engineer to do the same.

    - The analyst shall then assess its significance to the performance of tests and decide on the maintenance and calibrations required and subsequently, appoint a member of staff to have overall responsibility for the item of equipment.

  - Equipment Maintenance and Operation

8.3.1	Staff shall be assigned to have overall responsibility for certain items of equipment directly under the control of the laboratory. A list of staff assigned responsibility for such equipment shall be documented by the analyst and posted in the laboratory notice board.

8.3.2	The list will contain:

- File Number (If different from the equipment inventory number)

- Equipment Description

- Equipment Inventory number

- Person responsible

8.3.3	A list of all equipment will be generated and a calibration/maintenance form will be prepared for all equipment identified in above for each laboratory by the analyst and a copy will be forwarded to relevant maintenance authority which will contain weekly, monthly or annual calibration/maintenance schedule.

NOTE 1: The maintenance and calibration schedule may be made out as one, see example in Appendix C.

NOTE 2: It is advisable to schedule maintenance before calibration.

    - Newly acquired equipment and reference materials shall immediately be included in the maintenance schedule and the schedules shall be made available in the laboratory notice boards.

    - All equipment shall be maintained according to the maintenance schedule and the staff with overall responsibility for each item of equipment will ensure that such maintenance is carried out according to the manufacturer's recommendation.

    - When any equipment is due for maintenance, the officer responsible for it notifies the analyst for subsequent communication to the Technical Manager on maintenance and calibration.

    - The staff responsible for the equipment then forwards the equipment (if portable) and the request for calibration/maintenance as indicated in the calibration/maintenance form to the Technical Manager for subsequent communication to the appropriate maintenance authority.

    - If the equipment is non-portable the officer responsible for the equipment delivers the form to the Technical Manager.

    - On completion of maintenance the equipment is brought back to the lab if portable. The analyst will ensure that the equipment has been adequately serviced.

    - If an equipment breaks down or develops a fault in the course of its operation, the analyst or deputy will ensure that such an item of equipment that has suffered damage and has not been repaired is marked "OUT OF ORDER" and if not too bulky and space permits, it is removed from the general working area.

    - Apart from an annual equipment maintenance program, some equipment may require frequent intermediate checks in order to maintain confidence in their performance and/or calibration status. These checks will be defined, documented and carried out by the laboratory where appropriate.

  - Equipment Calibration

    - The analyst will then develop an annual calibration program or schedule for equipment maintenance identified in 8.1.1 above.

    - The analyst shall identify all equipment that requires calibration or as stated in the manufacturer's manuals and references.

    - All equipment shall be calibrated according to the calibration schedule issued and the staff with overall responsibility for each item of equipment shall ensure that the equipment is calibrated.

    - When any equipment is due for calibration, the officer responsible for it fills out the form in Appendix C and submits it to the Technical Manager for authorization.

    - The staff responsible for the equipment then forwards the equipment (if portable) and the calibration/maintenance form to the analyst. Same will happen if the equipment is not portable.

    - On completion of calibration, the equipment is brought back to the lab if portable. The analyst receiving the equipment then checks the equipment if it is in working order and records any observations, and/or whether the equipment has been brought back with the certificate. If such equipment is forwarded without calibration certificates, then the analyst shall place the equipment in a designated area and marked "ON HOLD", as they await the calibration certificates. In the case of equipment being calibrated in place, the officer responsible for the equipment carries out a similar check, as soon as the calibration is complete.

    - Once calibration certificates are ready they will be submitted to the Technical Manager by the contracted Metrology lab.

    - The Technical Manager checks the calibration certificate noting any errors indicated and decides whether or not the equipment is suitable for the intended use. This decision will be based on the specifications in the test methods available in the laboratory.

    - The Technical Manger shall extract the relevant information (errors, uncertainties) and make it readily available to users of the equipment.

    - If the said equipment is unsuitable for use it will be taken for repair and/or recalibration before it is put back to use.

    - If satisfactory, the calibrated equipment shall bear calibration stickers on them or on their packaging. Indicating, the date calibrated, next calibration due date, sticker serial number and calibration certificate number.

    - The calibration certificate is then filed in the calibration certificates file or the equipment file maintained by the Laboratory head of section.

    - Equipment overdue for calibration shall be considered to be non-conforming, thus the non- conformities/anomalies form shall then be filled followed by appropriate' action. However, for equipment with records of daily checks, e.g. balances, decision concerning their suitability may be easily obtained and the equipment may not necessarily be out of calibration.

    - Equipment operating on some ranges only shall be clearly labeled by use of a sticker on the equipment itself indicating the malfunctioning or the functioning ranges.

  - Computer-Based Equipment and Computers

    - The configuration management system maintained for all computer-based equipment and computers used in the test laboratories will include the following;

- A procedure for identifying and controlling the versions of hardware, software, firmware and documentation, both under installation and in use to ensure consistency, avoid errors and minimize the risk of unauthorized modifications;

- A facility for installation and re-installation or update/upgrade of computer systems; and

- A facility for safeguard (back-up) against loss and/or corruption during use and during storage. Backup shall be done in accordance with the appropriate procedures.

8.5.2 	In addition configuration management system covers the following:

- Application and system support software's

- Firmware and software library

- Data files and documentation files

- Memory, backing store, plug-in modules, board configuration and peripheral devices

  - Configuration Management System

    - General

      - Each computer-based equipment or computer systems shall be entered in to the laboratory inventory upon receipt and installation.

      - The analyst shall identify the computer system as indicated 8.1 above using the equipment numbers and an ERF Form filled out. Note: only one ERF shall be filled out for each computer system, and the identification of other parts of the system shall be as indicated in Appendix B.

      - The description of the computer system shall further be documented in Configuration Management Record (CMR). This includes: detailed list of all the individual parts of the system as identified, the manufacturer's identifications, including hardware and software.

      - The records of the current configuration, including hardware components, firmware versions and software installed shall be well maintained alongside the ERC open for the computer system. Note: initially, it consists of the supplier delivery and installation records where no changes are made.

      - The laboratory shall then develop instructions for use and overall description of the function of the system and where appropriate, a diagram showing the function of the data flow between the different parts of the system and the software employed.

  - Change Control and Version Control

    - Configuration changes shall be incorporated in the system only when one of the following activities has taken place:

i.       New purchased systems - during validation and subsequent handover to the laboratory;

ii.      Parts of the hardware of the system have been repaired;

iii. 	New hardware has been installed;

iv. 	A new version of the present software has been installed;

v. 	New software has been installed;

vi. 	New auxiliary equipment has been installed;

vii.    Re-installation of the software has been done;

viii.   Addition of or modification to software, automation facilities, etc.

    - All changes shall be identified in the CMR sheet.

    - The form will be submitted for assessment to Information and Communication Technology (ICT) department for recommendation and approval.

    - ICT department will provide or appoint an Information technology expert or engineer or the user of the system to implement the changes in liaison with the analyst and if necessary take affected equipment out of service and attach an "OUT OF USE" label. Only the authorized changes requested should be made and if further changes are necessary, they shall be added to the description of change and the form re-approved.

    - If the changes made affect the validation documentation and operating instructions, they shall be updated and referenced on the form. The program disks and any other copies of superseded files shall be withdrawn and the change signed off and logged in ERC (Equipment Record Change) sheet and the system brought back to service.

  - Validation and verification systems

    - The software of computer-based equipment shall be fully documented and validated before use. For the purposes of validation of computers used in test laboratories, it is sufficient to assume correct operation when input with known parameters, and also, the validation of individual components plus overall check on the dialogue between individual components and the controlling computer.

    - Acceptance test and Initial verification - the function of the system shall be verified and documented through the use of the following procedures as might be relevant:

i.   By using software test tools whenever these are available;

ii.   By simulating a test process through input of theoretical values into the computer unit of the system;

iii. By carrying out test/measurements on one or more objects with known outcome, e.g. use of certified reference materials in the case of chemical testing;

- By using certified test data sets where appropriate or available.

    - Continuous maintenance of the system - Verification of the system shall be done at regular intervals as scheduled by the laboratory and/or as part of the equipment maintenance program. Any recommendations made by the manufacturer shall be taken into consideration. In addition, access controls and protection from viruses shall also be included in the program, a period not exceeding one month is recommended. The analyst shall document procedures or instructions to be used for particular systems. All maintenance, either preventive or corrective shall only be performed by authorized personnel.

    - Specific Verifications

- Computer-based equipment with word-processing packages verification of such software will be achieved by manually checking the document hard copies.

ii. 	        For Data handling or signal processing systems and integration systems may not be validated in isolation from the instrument producing the original signal. This will therefore be validated in one go, by using reference standard materials and such validation will be acceptable if it gives correct results.

- Electronic transfer of data through the local Area Network (LAN) will be checked to ensure that no corruption occurs during transmission. Data transferred compact disc (CD) and flash disc will be similarly checked.

  - Reporting of Errors

Each computer system shall have a record book for documenting usage and checklist for daily system performance checks. If any errors are reported by the system the user shall record the error including the program in which it happened. The user shall initiate corrective action.

  - Storage of software items and system backup

    - A library shall exist to act as a store for software items, manuals and their configuration details. This shall be under the analyst.

    - All software items shall be deposited and signed for by the analyst. Only copies of this software may be kept in the laboratory.

    - All withdrawn software and obsolete items shall be archived. The software shall be kept free from dust, moisture and extremes of temperature. All software and subsequent updates will be authorized before use, having been first screened for viruses and validated.

    - The analyst will prohibit the use of unauthorized software on computers. Authorizations for use of obsolete versions of software will be rescinded to prevent further 'use. Such obsolete versions of software will be forwarded for safekeeping and archiving.

    - Backup of data files that form part of test records, shall be done in accordance with procedures.

  - Equipment Manuals

    - All equipment manuals, whether Operating manuals, Service and/or maintenance manuals, Installation Instructions or repair manuals shall be filed as stated in clause 8.1.3 of this procedure. If the manuals are bulky then they shall be kept in a separate file alphanumerically by manufacturer name or if the laboratory has adopted one equipment file system then, they will be with the other equipment records (i.e. calibration and maintenance).

    - The manuals shall be controlled by the equipment inventory number. That is, equipment inventory's number as given in appendix B and title of use - Operating manuals for XYZ equipment.

    - If a manual is to be used in the laboratory, say Operating instructions manual, an indication of the pages in which the operating instruction is found shall be made on the cover page bearing the "equipment number, i.e. "The Instructions given in pages X to Y are used verbatim for operations referred to by Test method ABC". This shall be signed and dated by the analyst.

    - If a new manual is received from the manufacturer, the old manual is withdrawn and taken for archiving as stated in procedure.

  - Equipment Movement

    - If equipment is transferred for any other reason other than for calibration and/or maintenance, the analyst will be notified on any necessary changes to be made.

  - Laboratory consumables:

    - An inventory of chemicals single use/disposable used that is significant for laboratory test and/or measurements will be maintained in the laboratory.

    - The laboratory will document instructions for cleaning such glassware and in a manner consistent with the testing protocol/methods to be applied.

  - APPENDIX:

9.1 Appendix A – Equipment inventory data sheet.

9.2 Appendix B – Maintenance workbook.

9.3 Appendix C - Calibration schedule.

9.4 Appendix D - Distribution list of controlled documents.

9.1 APPENDIX A: EQUIPMENT INVENTORY DATA SHEET

| S/No. | EQUIP LAB NO | EQUIP MODEL | EQUIP SERIAL | LOCATION | ANALYST |
| --- | --- | --- | --- | --- | --- |

  - AQC.QP2-F6

  - APPENDIX B: MAINTENACE WORKBOOK.

| S/No. | EQUIP LAB NO | EQUIP MODEL | EQUIP SERIAL | DATE OF SERVICE | DUE DATE |
| --- | --- | --- | --- | --- | --- |

AQC.QP2-F7

  - APPENDIX C: CALIBRATION SCHEDULE

| S/No. | EQUIP LAB NO | EQUIP MODEL | EQUIP SERIAL | DATE OF CALIBRATION | DUE DATE |
| --- | --- | --- | --- | --- | --- |

AQC.QP2-F8

9.4 APPENDIX D:  – DISTRIBUTION LIST OF CONTROLLED DOCUMENTS

| DOC. COPY NO. | DOC. VERSION |  |
| --- | --- | --- |
| 01 | 01 | QUALITY MANAGER |
| 02 | 01 | TECHNICAL MANAGER |
| 03 | 01 | ACCREDITATION BODY (KENAS) |

-End of document-